1.0 OBJECTIVE

1.1 To describe the procedure for operation of Agarose gel electrophoresis

2.0 SCOPE

2.1 This procedure is applicable for the operation of the Agarose gel electrophoresis unit in the R&D lab located at Oncosimis Biotech Pvt. Ltd. Plot no 3, IDA Uppal, Genpact Rd, Habsiguda, Hyderabad. TS - 500007

3.0 MATERIAL & SAFETY EQUIPMENT

- TAE or TBE

- EtBr

- Ladder

- 6x Loading Dye

Chemically Hazardous Ethidium Bromide (EtBr) is a potent mutagen and should be handled with extreme caution when mixed in the gel.

Safer alternative reagents include SYBR Safe, Gel Red, etc.

Electrical Hazards - The weighing balance, microwave, gel electrophoresis, UV transilluminator, etc. can lead to the electrical hazard in the laboratory. Electrophoresis units operating at 100 volts can provide a lethal shock.

Thermal Hazards - Laboratory personnel may be exposed to thermal hazards when heating agarose solutions.

Ultraviolet Radiation Hazards - Ultraviolet (UV) light boxes are often used in visualizing EtBr gels and pose potential exposures to UV radiation.

Safety training

All personnel (staff and students) handling hazardous chemicals and biohazard materials are required to complete the relevant safety training in Chemical Safety and Biosafety for BSL-2 Laboratories.

Personal protective equipment

Wear a long-sleeved lab coat, safety glasses/goggles, nitrile gloves (or glove material impermeable and resistant to the substance), and covered shoes. Wear insulated gloves if necessary

4.0 RESPONSIBILITY

4.1 Research Associate:

Shall operate the Agarose gel electrophoresis as per the procedure.

4.2 Senior Research Associate

Ensures that the Agarose gel electrophoresis operated as per the procedure.

4.3 R & D Head:

Ensures that the SOP is being followed as per the written instructions.

5.0 PROCEDURE

5.1 Pre-Operation

- Identified the location where ethidium bromide is used as "Designated Areas".

- The cytotoxic hazard symbol should be visible in the work area.

- Keep equipment clear of unintentional grounding points and conductors (e.g., sinks or other water sources, metal plates, aluminium foil, pipes, or other electrical equipment).

- Inspect equipment to be used and ensure all are in proper working condition. Report any equipment deficiencies before use.

5.2 During Operation

- Make sure that the power is off before connecting the electrical leads

- Connect the leads with only dried glove hands

- Do not run equipment unattended

- Never touch any part of the apparatus while the power is “ON”.

- If the electrophoresis buffer is spilled or is leaking from the gel tank, stop the run, turn off the equipment, clean up the benchtop and inspect the device immediately before proceeding

- Switch off the power completely and unplug the leads before opening the gel chamber lid or reaching inside the gel chamber

5.3 Preparing Agarose Gel

- Prepare agarose gel according to the instructions. Typically 0.7-2% agarose with TAE or TBE buffer is made. The volume of the gel will depend on the size of the casting tray

- Heat the agarose solution in a microwave until it starts bubbling and the agarose is completely dissolved. Do not over-boil. Ensure that the vessel containing the agarose has NO LID. Do not walk away from the heating agarose.

- Ensure the agarose powder is fully dissolved in the buffer. Let the hot gel preps cool to 50°-60°C before adding Ethidium bromide (EtBr) or its alternatives. (Note: Do not boil the agarose gel with EtBr or its alternatives)

5.4 Pouring of Agarose Gel into Gel Tank

- Pour the melted solution into a gel tray and allow it to cool until the gel is completely set. The gel should be between 3-5 mm thick. Check that no air bubbles are under or between the teeth of the comb

- When the gel has set solid (20-30 min at room temp), Remove the comb vertically out of the gel

- Place the agarose gel together with the gel tray into the electrophoresis chamber. Fill the reservoir with TAE or TBE buffer until the buffer covers the agarose gel.

5.5 Preparing and Loading the DNA Samples

- Mix the DNA samples with the desired amount of 6x gel-loading buffer. The maximum amount of DNA that can be applied to a slot depends on the number of fragments in the sample and their sizes

- Slowly load the sample mixture into the slots of the submerged gel using a disposable micropipette. Load ladder into slots on both the right and left sides of the gel. Avoid forming bubbles in the pipette tip when loading

5.6 Running gel

- Ensure the terminals are connected in the right way (red to red, black to black) and the samples are at the negative end of the gel, running toward the positive end

- Put a lid on the gel tank. Run the gel at an appropriate voltage (depends on the size of gel). Ensure current is flowing

5.7 Gel Documentation

- When the DNA samples or dyes have migrated a sufficient distance through the gel, turn off the electric current and remove the leads and the lid from the gel tank

- Transfer the gel to a small plastic box. Put the gel onto the UV transilluminator, close the cabinet door, and switch on the UV light. Adjust zoom and focus, then capture the image

6.0 REFERENCE

NIL

7.0 ABBREVIATIONS

TAE- Tris Acetate EDTA Buffer

TBE – Tris Borate EDTA Buffer

EtBr – Ethidium Bromide

8.0 ANNEXURES

8.1 Annexure – 1, Instrument Log Book

9.0 SUMMARY OF CHANGES

| Change control No | Reasons for revision |
| --- | --- |
| NA | NEW |