- 0 OBJECTIVE

  - To describe a procedure for cleaning of a Biostat B fermenter.

- 0 SCOPE

2.1 This procedure is applicable to the cleaning of Biostat B fermenter, located in                                   R&D lab at Oncosimis Biotech Pvt. Ltd. Plot no 3, IDA Uppal, Genpact Rd, Habsiguda,

Hyderabad.

- REFERENCE

- NIL

- MATERIAL & SAFETY EQUIPMENT

Materials:

- 3M Potassium Chloride (KCl) solution

- Autoclave

- Labolene solution

- Scrubbing Brush

Equipment:

- Name: Fermenter.

- Make: Biostat B

- Number

- RESPONSIBILITY

- Research Associate :

Shall operate the fermenter per the procedure

- Senior Research Associate:

Ensures that the fermenter is operated as per the procedure

- R & D Head :

Ensures that the SOI is being followed as per the written instructions

- PROCEDURE

| 6.1 On completion of the fermentation process, harvest the fermented broth |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 6.2 Close all the utility supply valves to the fermenter |  |  |  |  |  |  |  |  |
| 6.3 Disconnect the chilled water lines of the jacket heat exchanger & exhaust condenser |  |  |  |  |  |  |  |  |
| 6.4 Remove the DC motor from the vessel assembly |  |  |  |  |  |  |  |  |
| 6.5 Remove the temperature sensor, rinse with water, wipe dry, and store |  |  |  |  |  |  |  |  |
| 6.6 Remove the cables of pH and D.O. probes and fit the probe ends with protective caps |  |  |  |  |  |  |  |  |
| 6.7 Remove the foam probe cables and store them in a clean place |  |  |  |  |  |  |  |  |
| 6.8 Remove the sampler |  |  |  |  |  |  |  |  |
| 6.9 Take out the silicone tubing end from the feed medium/solutions, start the feed pumps to empty the hold-up volume in the tubings into the fermenter |  |  |  |  |  |  |  |  |
| 6.10 Cover the silicone tubing ends with parchment paper and secure them with rubber bands |  |  |  |  |  |  |  |  |
| 6.11 Close the inlet air filter tubing with a pinchcock and cover the end with parchment paper |  |  |  |  |  |  |  |  |
| 6.12 Close the silicone tubing on the harvest line with a pinchcock and cover its end with parchment paper |  |  |  |  |  |  |  |  |
| 6.13 Fill the fermenter with 5lt. of purified water |  |  |  |  |  |  |  |  |
| 6.14 Autoclave the fermenter |  |  |  |  |  |  |  |  |
| 6.15 After the fermenter has cooled down to room temperature, remove the following: - |  |  |  |  |  |  |  |  |
| D.O. probe: Rinse thoroughly with distilled water and store in a clean dry place |  |  |  |  |  |  |  |  |
| pH probe: Rinse thoroughly with distilled water and store in a clean dry place |  |  |  |  |  |  |  |  |
| Air filters: Store in a clean place |  |  |  |  |  |  |  |  |
| Silicone Tubings: Wash and store in a clean place |  |  |  |  |  |  |  |  |
| 6.16 Unscrew the bolts and remove the top plate |  |  |  |  |  |  |  |  |
| 6.17 Decant the 5lt. water from the fermenter and drain it to the kill tank |  |  |  |  |  |  |  |  |
| 6.18 Rinse the vessel with purified water and drain the rinse to the kill tank |  |  |  |  |  |  |  |  |
| 6.19 Wash the vessel, top plate, and other accessories. |  |  |  |  |  |  |  |  |
| 6.20 After washing, leave the vessel on a clean surface for drying |  |  |  |  |  |  |  |  |
| Precautions: Handle the pH probe and glass vessel with care during washing, to prevent breakage |  |  |  |  |  |  |  |  |

- ABBREVIATIONS

NIL

- ANNEXURES

NIL

9.0 SUMMARY OF CHANGES

| Change control No | Reasons for revision |
| --- | --- |
| NA | NEW |