1.0 OBJECTIVE

1.1 To describe the standard operating procedure for FPLC– AKTA purifier system

2.0 SCOPE

2.1 This procedure is applicable for operation of FPLC AKTA purifier system located in the R&D lab at Oncosimis Biotech Pvt. Ltd., Plot no 3, IDA Uppal, Genpact Rd, Habsiguda, Hyderabad. TS - 500007

3.0 MATERIAL & SAFETY EQUIPMENT

- EtOH 20% (HPLC grade) – minimum 1L

- MilliQ – minimum 1L

- Appropriate column

- Appropriate buffers chilled in AKTA refrigerator

4.0 RESPONSIBILITY

4.1 Research Associate:

Shall operate the FPLC standard operating procedure – AKTA purifier system as per the procedure

4.2 Senior Research Associate:

Ensures that the Gel FPLC standard operating procedure for the AKTA purifier system is operated as per the procedure.

4.3 R & D Head:

Ensures that the SOP is being followed as per the written instructions

5.0 PROCEDURE

5.1 Pre-run checklist

- Check that the main waste bottle is empty (lower refrigerator chamber)

- Check that the pump waste bottle is empty (placed at the left side of the FPLC with two waste tubes).

- Check that the fraction collector base unit status light is green colored and that the carousel is positioned securely in its place

- Check that there are enough clean tubes in fraction rack, with the “A” row

- Check there is sufficient volumes of buffers, MilliQ and EtOH 20% for the coming run(s)

- Check pump and tubing: Tube A1 should be placed in MilliQ bottle; Tube A1 should be connected to position #1 on buffer selection valve IV-908, In case there is need to    transfer any tubing from one bottle to another, make sure you wash the tube with MilliQ water before placing it in its new position

5.2 Preparing the system for a run

Note:

Switch on PC and initiate the “Unicorn” software package (see desktop icon); In order to simultaneous execute several commands at the control panel, press “insert” and at the end of the command listing press “execute”. Following which, four windows will open.

- Login info ⇒ User: default, Password: default

- In the Unicorn software, prompt the system control window;

- Go to Manual in the menu items (or press Ctrl + M).

PC monitor checklist:

- Choose Flow path:

- Check that “injection valve” is set to load

- Column position ⇒ check that it is set at “Position1 bypass”

- Buffer ValveA1 ⇒ Check that the pump inlet is set to A11

- Pump inlet ⇒ A1

- Choose “Pump” and check the following before continuing:

- Flow rate ⇒ should be set according to the permitted flow rate of the column (consult column’s accompanying leaflet for the recommended flow rate)

- Gradient ⇒ Check that inlet %B is set to “0”.

- Alarm & monitor ⇒ set alarm pressure to 0.25MPa

Wash system’s A1 pump:

- Go to “pump” ⇒ “Pump Wash Explorer”, Choose “Inlet A11”. In this automatic program, the pump is washed for 4 minutes with MilliQ pumped from inlet A11.

- If pump B1 is to be used, wash pump B1 with MilliQ (wash and transfer tubing B11 to MilliQ bottle)

- Wash B11 tubing with an appropriate buffer (if it is MonoQ it is usually high salt buffer).

5.3 Wash inlet A12 tubing with MilliQ:

- Verify that the system is at standby (check to see that flow rate is “0.00”).

- Rinse A12 tubing end thoroughly with MilliQ and place in MilliQ bottle

- Change flow path⇒” Buffer Valve A1” ⇒ “A12”

- Set flow rate ⇒ 10ml/min

- Wash in 20ml of MilliQ

5.4 Preparation of loop for small volume injection (up to 5ml):

- Connect loop to injection valve via ports 2 and 6

- Wash loop with MilliQ - Fill 10ml syringe with MilliQ, connect the syringe to the injection port (screw-type) and pump the MilliQ while monitoring the outlet tubing drizzle (if it is not constant there might be air bubbles in the system; continue pumping till drizzle is constant).

- Repeat wash with the equilibration buffer intended to use for protein purification (according to the column used)

- Injection valve is ready for a run

5.5 Connect column:

- Remove the plastic capping from the column ends and mount the column onto the column arms

- If it is a small column, it is possible to connect the column directly to pump valve#3 at any available position

- Connect column’s lower part to pump valve#3 (upper left module) via tubing on position# 8

- Connect the column’s upper port to pump valve #2 via position #8 in a drop-to-drop style (this step ensures that no air bubbles will enter the column)

- Go to Flow path ⇒ “column position” ⇒ “A18”

- Check Flow path ⇒ “Buffer Valve A1” ⇒ Choose appropriate tubing

- Go to Flow rate ⇒ 0.5ml/min (make sure the system is at “Run” mode)

- The inlet tube edge should start to drizzle – fill the tubing socket with liquid and only then screw in the connecter while pressing down.

- Monitor the curve on screen and see that the curves stabilize after several minutes.

5.6 Wash column with protein’s buffer2:

- Verify that the system is at standby (check to see that flow rate is “0.00”).

- Transfer A12 tubing to buffer’s bottle

- Set flow rate ⇒ 0.5ml/min

- Wash the column with two-column volumes with the buffer

- The column is ready for the sample run

5.7 Purging the air from pump

- Connect purge tubing with a syringe to the purge valve

- Unscrew purge valve half a turn and connect purge and draw approximately 10ml or until no air bubbles can be seen in the pump tubing and the purge tubing.

- Close the purge valve and disconnect the purge tubing

5.8 Sample run – Loading via buffer selection valve (IV-908 module)

Note: This injection method is suitable for >5ml protein solutions; lower volume

should be injected via loop (see detailed procedure below)

- Make sure the relevant tubings are connected to the appropriate buffers and the buffer levels are sufficient for the subsequent run

- Make sure the flow rate is set to “0 ml/min”, column position is set accordingly (usually at position 8)

- Transfer the injection tube into your sample tube (should be 50ml flacon) and screw the cap; make sure the tube’s ending is at the lower part of the tube (this position should be maintained along the whole length of the injection step).

- Set Flow path ⇒ Buffer Valve A1 ⇒ “A18”

- Initiate injection by setting the flow rate to the column manufacturer’s recommendation; it is advisable to use a 1ml/min flow rate

- Monitor the volume of sample; when 2-3ml is left in the tube stop injection by setting the flow rate to “0ml/min”

- (optional – IEC application):

  - Go to flow path ⇒ Set “A12”

  - Go to pump ⇒ gradient ⇒ set “%B” as 40% for 40’ at 1ml/min

  - Prepare for sample collection:

    - Go to “Frac” ⇒ main fractionation

    - Choose frac size ⇒ “10ml”, “Row-by-Row”

Note: Don’t press execute until the sample starts to elute!

    - Once there is a need to move to the next tube, press “Feed tube”

    - Fractionation stop is initiated via “Fractionation stop” command ⇒ notice that flow path is switched to “Waste”.

5.9 Sample run – Loading via loop (INV-908)

Before starting the procedure, remove aggregates from the protein solution by either centrifuging sample at >13Krpm for a minimum of 10 mins at 4°C or utilize 0.22μM filter (which will be connected directly to the injection port)

- Connect loop via ports 3 and 6 on INV-908

- Wash loop with MilliQ, at least x2 loop volume

- Repeat step 2 with equilibration buffer used for protein purification

- Connect the needle to the appropriate volume syringe and wash the needle with one syringe volume of MilliQ and one syringe volume of equilibration buffer

- Draw protein sample into a syringe to remove any bubble formed

Note: Before injection, make sure flow rate is correctly set, the flow path is set to column position, buffer Valve choice is set to the correct tubing, and the injection valve is set to “load”.

- If using a filter, connect the filter to a syringe and connect the filter nozzle to the injection port

- Inject your sample at half the loop volume and set injection status to “Inject”

- Wait at least x2 loop volume and then re-set the injection valve to “Load”

- Monitor progression and collect the samples

6.0 REFERENCE

NIL

7.0 ABBREVIATIONS

NIL

8.0 ANNEXURES

8.1 Annexure – 1, Instrument Log Book

9.0 SUMMARY OF CHANGES

| Change control No | Reasons for revision |
| --- | --- |
| NA | NEW |

END OF DOCUMENT