1.0 OBJECTIVE

1.1 To describe the procedure for the operation of SDS-PAGE

2.0 SCOPE

2.1 This procedure is applicable for the operation of SDS-PAGE unit located in the R&D lab at Oncosimis Biotech private limited Plot no 3, IDA Uppal, Genpact Rd, Habsiguda, Hyderabad, TS

3.0 MATERIAL & SAFETY EQUIPMENT

- Ladder

- 4x Loading Dye

Electrical Hazards: The weighing balance, microwave, gel electrophoresis, UV transilluminator, etc. can lead to an electrical hazard in the laboratory. Electrophoresis units operating at 100 volts can provide a lethal shock

4.0 RESPONSIBILITY

4.1 Research Associate:

Shall operate the SDS-PAGE as per the procedure

4.2 Senior Research Associate:

Ensures that the SDS-PAGE operated as per the procedure.

4.3 R & D Head:

Ensures that the SOP is being followed as per written instructions

5.0 PROCEDURE

5.1 Gel Cassette Preparation Hand cast Gels:

a. Place the casting frame upright with the pressure cams in the open position and facing          forward on a flat surface

b. Select a spacer plate of the desired gel thickness and place a short plate on top of it

c. Orient the spacer plate so that the labelling is up. Slide the two glass plates into the casting frame, keeping the short plate facing the front of the frame (side with pressure cams)

d. When the glass plates are in place, engage the pressure cams to secure the glass cassette sandwich in the casting frame, check that both plates are flushed at the bottom.

e. Place the casting frame into the casting stand by positioning the casting frame (with the locked pressure cams facing out) onto the casting gasket while engaging the spring-loaded lever of the casting stand onto the spacer plate.

Note: The grey casting stand gaskets must be clean and dry. The casting stand gaskets are made of a special thermoplastic material that swells when soaked in water, so we recommend that you do not soak the gaskets for prolonged periods prior to casting. If the

gaskets do get accidentally soaked and display swelling and/or deformation, just allow them to air dry and they will regain their original shape, size and performance.

5.2 Electrophoresis Module Assembly and Sample Loading

a. Set the clamping frame to the open position on a clean flat surface.

b. Place the first gel sandwich or gel cassette (with the short plate facing inward) onto the gel supports; gel supports are molded into the bottom of the clamping frame assembly; there are two supports on each side of the assembly.

c. Using one hand, gently pull both gels towards each other, making sure that they rest firmly and squarely against the green gaskets that are built into the clamping frame; make certain that the short plates sit just below the notch at the top of the green gasket

d. While gently squeezing the gel sandwiches or cassettes against the green gaskets with one hand (keeping constant pressure and both gels firmly held in place), slide the green arms of the clamping frame over the gels, locking them into place. Alternatively, you may choose to pick up the entire assembly with both hands, making sure that the gels do not shift, and simultaneously sliding both arms of the clamping frame into place. The arms of the clamping frame push the short plates of each gel cassette up against the notch in the green gasket, creating a leak-proof seal (check again to make certain that the short plates sit just below the notch at the top of the green gasket). At this point, the sample wells can be washed out with running buffer.

5.3 Sample Loading

a. Fill the assembly (upper chamber) with buffer to just under the edge of the outer gel plate.

b. Load the samples into each of the assemblies while they are sitting on a flat surface, outside of the tank

c. Load the samples into the wells with a Hamilton syringe or a pipet using gel loading tips

d. If using Bio-Rad’s patented sample loading guide, place it between the two gels in the electrode assembly. Sample loading guides are available for 9, 10, 12, and 15-well formats

e. Use the sample loading guide to locate the sample wells. Insert the Hamilton syringe or pipet tip into the slots of the guide and fill the corresponding wells

Note: Load samples slowly to allow them to settle evenly on the bottom of the well. Be careful not to puncture the bottom of the well with the syringe needle or pipet. Note: Samples may be loaded in the modules before placing the modules into the tank. Samples may also be loaded in the modules after the modules have been placed into the tank. Both methods will produce acceptable results. In both instances, the assembly (upper chamber) and the tank (lower chamber) should be filled with buffer according to the instructions

5.4 Placement of the Electrode Assemblies in the Mini-PROTEAN:

Note: required total buffer volume, 700 ml for 2 gels; 1000 ml for 4 gels. The Mini-PROTEAN the electrode assembly (back position) and the companion running module (front position)

a. Begin by placing the tank on a flat surface, with the front of the tank facing you (the front of the tank is the face that has the 2-Gels and 4-Gels line markings); when oriented properly, the red marking on the top inside edge of the tank will be on your right, and the black marking on the top inside edge of the tank will be on your left.

b. If running 2 gels only, you will be using just the electrode assembly, so place this assembly in the back position of the cell, making sure that the red (+) electrode jack matches the red marking on the top right inside edge of the tank. c. If running 4 gels, place the electrode assembly, in the back position, and the companion running module (no banana plugs) in the front position. Make sure that in both instances the red (+) electrode is matching with the red marking on the top inside the right edge of the tank.

c. Note that incorrect orientation will not permit proper placement of the lid

5.5 Gel Removal

a. After the electrophoresis is complete, turn off the power supply and disconnect the electrical leads.

b. Remove the tank lid and carefully lift out the electrode assemblies. Pour off and discard the running buffer.

Note: Always pour off the buffer before opening the arms of the assembly, to avoid spilling the buffer.

c. Open the arms of the assembly and remove the gel cassettes.

d. Remove the gels from the gel cassette by gently separating the two plates of the gel cassette.

e. Remove the gel by floating it off the plate by inverting the gel and plate under fixative or transfer solution, agitating gently until the gel separates from the plate.

f. Rinse the Mini-PROTEAN electrode assembly, clamping frame, and mini tank with distilled, deionized water after use

6.0 REFERENCE

NIL

7.0 ABBREVIATIONS

TAE- Tris Acetate Buffer

EtBr – Ethidium Bromide

8.0 ANNEXURES

8.1 Annexure – 1, Instrument Log Book

9.0 SUMMARY OF CHANGES

| Change control No | Reasons for revision |
| --- | --- |
| NA | NEW |

END OF DOCUMENT