- 0 OBJECTIVE

  - To describe the procedure for the operation of Tangential Flow Filtration

- 0 SCOPE

2.1 This procedure is applicable for the operation of Tangential Flow Filtration, located

in the R&D lab at Oncosimis Biotech Pvt. Ltd., Plot no 3, IDA Uppal, Genpact Rd, Habsiguda, Hyderabad. TS, India 500007

- REFERENCE

- NIL

- MATERIAL & SAFETY EQUIPMENT

Materials

- Water

- 0.5 N NaOH

- Peristaltic Pump, Tubings

Equipment

- Name: Tangential Flow Filtration.

- Make: Soleus

- Number

- RESPONSIBILITY

- Research Associate:

Shall operate the Tangential Flow Filtration as per the procedure

- Senior Research Associate:

Ensures that the Tangential Flow Filtration is operated as per the procedure

- R & D Head:

Ensures that the SOP is being followed as per written instructions

- PROCEDURE

  - Operation

    - Connect the cassette to the holder

    - Attach master flex tubing id 18 to feed input, retentate, permeate hose

    - Supplement the filtration unit with a peristaltic pump which is compatible to hold feed input tubing

    - Switch on the peristaltic pump and press start on the machine to send the feed to the filtration unit

    - Maintain 0 bar pressure and ensure complete feed input comes out of the retentate on the first run.

    - Input the samples containing protein through feed input and apply pressure about 1 bar using the valve located at retentate.

    - Collect permeate in a separate beaker.

    - Upon completion, depressurize to 0 bar using the valve

    - Press the stop button in the peristaltic pump and switch of the machine

  - Sanitization

    - After completion of the operation, pass 2-3 liters WFI by maintaining the 0 bar pressure in both retentates and feed input regulating valve

    - During the initial passage of WFI, ensure retentate and permeate are collected in different beakers

    - Apply about 1 bar pressure using retentate regulating valve and pass 2-3  WFI and collect the permeate

    - Pass through 0.5 N NaOH about 2 liters through feed input and with about 1 bar pressure

    - Subsequently, pass enough WFI through feed input and check the pH of the permeate reach pH of WFI

    - Turn off the peristaltic pump and clip all the tubing

- ABBREVIATIONS

  - WFI 	- Water for Injection

  - NaOH 	- Sodium Hydroxide.

- ANNEXURES

  - Annexure – 1, Instrument Log Book.

- SUMMARY OF CHANGES

| Change control No | Reasons for revision |
| --- | --- |
| NA | NEW |

END OF DOCUMENT